|  | 1300 North 17th Street, Suite 1752 Arlington, VA 22209, USA +1-703- 841-3281 http://dicom.nema.org E-mail: dicom@medicalimaging.org |
| --- | --- |

Meeting MINUTES-Rev.1

###### DICOM WORKING GROUP SIX

(Base Standard)

June 11-14, 2012

### DIN, Berlin, Germany

| Members Present | Represented by |  | M | T | W | Th |
| --- | --- | --- | --- | --- | --- | --- |
| Agfa Healthcare | Robert Horn |  |  | P |  |  |
| DatCard Systems | Jere Darling |  | x | x | x | x |
| GE Healthcare | Andrei Leontiev |  |  | x | x | x |
| JIRA | Makoto Suzuki |  | x | x | x | x |
| Philips Healthcare | Bas Revet |  | x | x | x | x |
| Siemens Healthcare | Bjorn Nolte |  | x | x | x | x |
| Toshiba Medical Research Institute USA | Kevin O’Donnell |  |  | x | x | x |
| Members Absent | Voting Representative |  |  |  |  |  |
| CoreLab | David Clunie |  |  |  |  |  |
| * ICCAS (Innovation Center Computer Assisted Surgery (ICCAS)) | Burgert, Oliver |  |  |  |  |  |
| *Carl Zeiss Medical Software | Tobias Kurzke |  |  |  |  |  |
| *DeJarnette Research Systems | Fred Ganong |  |  |  |  |  |
| Konica Minolta Med. Imaging | Michael Laconti |  |  |  |  |  |
| Merge Healthcare | Ronald Cowan |  |  |  |  |  |
| *Societe Francaise de Radiologie | Joel Chabriais |  |  |  |  |  |

| Alternate Voting Representatives, Observers, Others Present | Affiliation |  | M | T | W | T |
| --- | --- | --- | --- | --- | --- | --- |
| Culp, Timothy | Wg-27-Harris Corp. |  |  |  | x | x |
| Kanzaki, Masatoshi | WG-11-Real Vision |  |  |  | P | P |
| Lemke, Heinz | WG-24-ICCAS |  |  | x |  |  |
| Liebmann, Philipp | WG-24-ICCAS |  | x | x |  |  |
| Matsui, Norihisa | WG-11-Shimadzu |  |  |  | P | P |
| Schadt, Christof | WG-07-Brainlab |  |  | x | x |  |
| Solomon, Harry | WG-26-GE Healthcare |  |  |  | P |  |
| Stephen Vastagh | MITA Staff |  | x | x | x | x |
| Whitby, Jonathan | Toshiba |  |  |  | x | x |

(* = Does not count toward a quorum,  P = Phone and/or web,  X = Present)

Presiding Officer:							Bas Revet, Chair

# Opening

- The Chair called the meeting to order at 09.30 on June 11.  Participants identified themselves and their employers.  The secretary reported that a quorum will be present, on June 12.  Members approved the agenda.  The minutes of the January 2012 meeting (Rev. 1) were reviewed. The minutes were approved with revising the Sup. 154 section to indicate approval for Public Comment.  The Secretary reviewed the antitrust rules.

# Secretary’s Report

Committee Secretary Stephen Vastagh reported on the ballot distributed following the March. 2012 meeting of WG-06.  The ballot closed 2012-05-25; the ballot met DICOM’s quorum requirements and passed.

| ## BALLOT Closed | ## APPROVE | ## DISAPPROVE | ## ABSTAIN | COMMENTS |
| --- | --- | --- | --- | --- |
| 2012-05-25 | - 23 | 0 | 6 | 3 |

# Decisions: on CPs in Letter Ballot CPack-66 (Correction Package No. 66) >  Comments > Final Text or other Disposition

- The CPs in the Letter Ballot CPack-66 are listed below together with discussion of comments and decision on Final Text or other disposition.

| CP | Name |  | Consideration of Comments and Other Discussion | Dis-position |
| --- | --- | --- | --- | --- |
| 1148 | Add Pathology Protocol Codes |  | Hologic and Siemens comments were accommodated. Need to check with D. Clunie on UID issues. No reference to TID, Bas will edit. | FT |
| 1176 | Change CID of Anatomic Region Sequence for X-Ray 3D Angiography |  | No comment. | FT |
| 1191 | Fix Typo in TAG for Lossy Image Compression Method |  | No comment. | FT |
| 1192 | Fix Typo RT Beams Delivery instruction module and dictionary |  | No comment. | FT |
| 1193 | Fix Typo in table C.8.27.2-4 |  | No comment. | FT |
| 1194 |  |  | Toshiba-KOD comment accepted. B. Revet will edit. | FT |

- WG-06 decided unanimously to approve, as final text, the CPs listed above.

# Decisions: on CPs that were in Voting Package for June 2012

- WG-06 reviewed and discussed the CPs that were in the VP (Voting Package) coming into this meeting.  The CPs were approved for inclusion in Letter Ballot (CPack 67) or directed for further work or review as shown below:

| CP# | Voting Package for June 2012 -- Title | Assigned to | Discussion | Disposition |
| --- | --- | --- | --- | --- |
| 739 | Add View Code Sequence to CT, MR, SC for Cardiac Views (moved to assigned for coordination with other WGs) | D. Clunie |  | LB |
| 1112 | Add SR Structure for ROI Statistics | H. Solomon | Still on hold. If no answer on SNOMED Code by week’s end, keep in VP | VP Aug |
| 1137 | Structure Set Relationship Sequence | U Busch |  | LB |
| 1138 | Average Beam Dose Parameter | U Busch |  | LB |
| 1180 | Use LOINC_Short Name | H. Solomon | Not ready for LB.. | VP Aug |
| 1195 |  | B Revet |  | LB |
| 1197 | Unspecified code should be forbidden in extensible context groups for mandatory content items | D Clunie |  | LB |
| 1198 | Clarify encoding of Surface Segmentation Point Indices | D Clunie |  | LB |
| 1199 | Add Modifier to Processing Parameters Sequence | D Clunie |  | LB |
| 1202 | Correct code for Liver | B Revet |  | LB |
| 1203 | PDR Pulse Details in RT Brachy Session Record | U Busch | Not ready for LB.. | VP Aug |
| 1204 | Identify specific Source used in RT Plan | U Busch |  | LB |
| 1205 | Device Parameters in CT Images | U Busch |  | LB |
| 1206 | Dose Summation Type Clarification | U Busch |  | LB |

# Letter Ballot CPack-67

- There were no additional CPs approved for letter ballot other than those from the June Voting Packet.

Accordingly, WG-06 decided unanimously to submit for legal review and thereafter circulate in a Letter Ballot CPack 67, for vote by the DICOM Standards Committee, the Correction Proposals (CPs) listed below:

| CP# | Letter Ballot to Be Issued CPack-67 – Title | Assigned to |  | Disposition |
| --- | --- | --- | --- | --- |
| 739 | Add View Code Sequence to CT, MR, SC for Cardiac Views (moved to assigned for coordination with other WGs) | D. Clunie |  | LB |
| 1137 | Structure Set Relationship Sequence | U Busch |  | LB |
| 1138 | Average Beam Dose Parameter | U Busch |  | LB |
| 1195 |  | B Revet |  | LB |
| 1197 | Unspecified code should be forbidden in extensible context groups for mandatory content items | D Clunie |  | LB |
| 1198 | Clarify encoding of Surface Segmentation Point Indices | D Clunie |  | LB |
| 1199 | Add Modifier to Processing Parameters Sequence | D Clunie |  | LB |
| 1202 | Correct code for Liver | B Revet |  | LB |
| 1204 | Identify specific Source used in RT Plan | U Busch |  | LB |
| 1205 | Device Parameters in CT Images | U Busch |  | LB |
| 1206 | Dose Summation Type Clarification | U Busch |  | LB |

# Discussed: Assigned CPs (New CPs in March 2012 and Other Previously Assigned CPs)

(The table below records discussion of assigned CPs OTHER THAN THOSE that were in Letter Ballot CPack 66 (recorded in Item/Table 3) or in the June Voting Packet (recorded in Item/Table 4.)

| CP# | Title | Assigned to | Discussion | Dis-position |
| --- | --- | --- | --- | --- |
| 1112 | Add SR Structure for ROI Statistics | H. Solomon | Still on hold. If no answer on SNOMED Code by week’s end, keep in VP | VP Aug |
| 1152 | Additional document titles for Key Object Selection Document | R. Horn |  | Aug VP |
| 1200 | Add attributes to Surface Mesh Module | U. Busch | Use cases were supplied | Aug VP |
| 1201 |  | U. Busch | Found grammer issues, revised, will confirm with H Blendinger if revisions are OK | Aug VP |
| 1210 | Add Modifier to Processing Parameters Sequence | D. Clunie |  | Aug VP |
| 1211 | Inconsistent Mean Point Distance in Surface Segmentation | D. Clunie |  | Aug VP |
| 1212 | SR IOD with SCOORD3D | D. Clunie |  | Aug VP |
| 1213 | Clarify homogeneous transformation matrix constraints | D. Clunie | Input received from WG-07. | Aug VP |
| 1214 | Add Tracking Identifier Template to Measurement Templates | D. Clunie |  | Aug VP |
| 1215 | Require GSPS for Enhanced PET Storage for consistency with other enhanced images | D. Clunie |  | Aug VP |
| 1216 | Instance Creation Date and Time not changed on coercion | D. Clunie |  | Aug VP |
| 1217 | Add Extensible SR SOP Class for new types of Content Item | D.. Clunie | Need discussion with D Clunie and R horn present in Aug | Remains assigned |
| 1218 | Correct further errors in UCUM cades and related meanings | B Nolte | B Revet reported that the PET community is in agreement | Aug VP |
| 1221 | Add note on VR conflict in data dictionary | D. Clunie | References a very old supplement. Cannot be fixed. Note is appropriate. | Aug VP |
| 1222 | Amend DIMSE Service status codes | B. Revet |  | Aug VP |
| 1223 |  | B. Revet |  | Aug VP |
| 1234 | Add GBK and GB2312 Character Sets for Chinese Text Encoding | R Horn | New CP in June from CIMICS | Aug VP |

# VP (Voting Package) for Aug 2012

- Based on discussions and decisions as recorded in Tables 3, 4, and/or 6, the Voting package for August 2012 is:

| CP# | Title | Assigned to | Discussion | Dis-position |
| --- | --- | --- | --- | --- |
| 1152 | Additional document titles for Key Object Selection Document | R. Horn |  | Aug VP |
| 1180 | Use LOINC_Short Name | H. Solomon |  | Aug VP |
| 1200 | Add attributes to Surface Mesh Module | U. Busch |  | Aug VP |
| 1201 |  | U. Busch |  | Aug VP |
| 1203 | PDR Pulse Details in RT Brachy Session Record | U. Busch |  | Aug VP |
| 1210 | Add Modifier to Processing Parameters Sequence | D. Clunie |  | Aug VP |
| 1211 | Inconsistent Mean Point Distance in Surface Segmentation | D. Clunie |  | Aug VP |
| 1212 | SR IOD with SCOORD3D | D. Clunie |  | Aug VP |
| 1213 | Clarify homogeneous transformation matrix constraints | D. Clunie |  | Aug VP |
| 1214 | Add Tracking Identifier Template to Measurement Templates | D. Clunie |  | Aug VP |
| 1215 | Require GSPS for Enhanced PET Storage for consistency with other enhanced images | D. Clunie |  | Aug VP |
| 1216 | Instance Creation Date and Time not changed on coercion | D. Clunie |  | Aug VP |
| 1218 | Correct further errors in UCUM cades and related meanings | B Nolte |  | Aug VP |
| 1221 | Add note on VR conflict in data dictionary | D. Clunie |  | Aug VP |
| 1222 | Amend DIMSE Service status codes | B. Revet |  | Aug VP |
| 1223 |  | B. Revet |  | Aug VP |
| 1234 | Add GBK and GB2312 Character Sets for Chinese Text Encoding | R Horn |  | Aug VP |

# CPs Remaining Assigned

8.1	CPs Discussed at this Meeting – Remaining Assigned

As recorded in Tables 4, 5, and/or 7:

| CP# | Title | Assigned to | Discussion | Dis-position |
| --- | --- | --- | --- | --- |
| 1217 | Add Extensible SR SOP Class for new types of Content Item | D.. Clunie | Need discussion with D Clunie and R horn present in Aug | Remains assigned |

8.2	CPs NOT Discussed at this Meeting – Remaining Assigned

The following previously assigned correction proposals were not discussed and dispatched as noted:

| CP# | Title | Assigned to | Discussion |  |
| --- | --- | --- | --- | --- |
| 145 | Curve Data Value Representation Clarification | D. Clunie |  |  |
| 235 | Clarify NM detector start angle usage | J. Pohlhammer |  |  |
| 296 | Eliminate Duplicate Terms and Provide Missing Definitions | H. Solomon |  |  |
| 375 | - Correct CID 4017 – Intra-oral Anatomy Modifier | K. O’Donnell |  |  |
| 430 | Additions to PS 3.16 for Cardiac Cath SR | H. Solomon |  |  |
| 450 | Correct Myocardium Mass | D. Sluis |  |  |
| 464 | Add SOP Common to the Print-related N-CREATEs | R. Horn |  |  |
| 471 | Add Common Finding Template | H. Solomon |  |  |
| 591 | Attribute Identifier List for Errors in Sequence Attributes | H. Solomon |  |  |
| 596 | Add Request and Schedule Information to Results and MPPS | R. Horn |  |  |
| 597 | Instance Availability Notification Structure | H. Solomon |  |  |
| 757 | Add RTSS ROI Feature UID | H. Solomon |  |  |
| 802 | Incorporate Fixes on 702 and 706 | D. Clunie |  |  |
| 804 | Indicator for Cone Beam CT | K. O’Donnell |  |  |
| 811 | Note for Forwarders Regarding SCP-assigned Instance UIDs during N-CREATE | D. Harvey |  |  |
| 812 | Clarification that Pixel Padding Value Is a Raw Value before Any Transformations | D. Harvey |  |  |
| 833 | Transfer Syntax Availability and Selection for Q/R | R. Horn |  |  |
| 836 | Clarify Usage of Attributes for Normalized IODs | K. O’Donnell |  |  |
| 838 | Consistent Naming of Error Code Tables | K. O’Donnell |  |  |
| 934 | Fix Inconsistency of C-FIND Only Number of Patient Related XXX Attributes | D. Harvey |  |  |
| 941 | More Clarification of Instance References | D. Clunie |  |  |
| 991 | Clarify Handling of Private Data in the Retrieve Without Bulk Data Service | R. Horn |  |  |
| 992 | Clarify Allowable Length Values for Certain PDUs | K. O’ Donnell |  |  |
| 1005 | Make Keywords Valid Java/C Identifiers | R. Horn |  |  |
| 1006 | Add Keywords to Command Dictionary | R. Horn |  |  |
| 1013 | RT Ion Beams Recording Clarifications | WG-07 |  |  |
| 1019 | Add Body Part at the Study Level | KO’Donnell |  |  |
| 1030 | Re-factor Common Attributes for Enhanced Images | B. Revet |  |  |
| 1031 | Use of OIDs and Non-ASCI Characters in Codes | R. Horn |  |  |
| 1032 | Fix Position Angle Sign Inconsistency in Mammography Image | K. O’Donnell |  |  |
| 1066 | Encoding of Attributes with Value Length > 64KiB with Explicit VR | R. Horn |  |  |
| 1086 | New DICOMDIR ZIP File Extension | R. Horn |  |  |
| 1106 | Assign a Namespace for Audit Schema Elements | R. Horn |  |  |
| 1115 | Add Dose Reference Point Codes for CR/DR | K. O’Donnell |  |  |
| 1126 | Corrections LOINC Codes Sup78 | B. Revet |  |  |
| 1127 | Add Fields for Organ Dose to Dose SR | K. O’Donnell |  |  |
| 1163 | Conformance for Application Hosting | H Solomon |  |  |
| 1164 | Coronary Artery Codes | H Solomon |  |  |
| 1173 | Clarify basis for Total Fluoro Time and Total Acquisition Time in XA Dose Report | K O’Donnell |  |  |
| 1189 | Add usage of Temporal Position Time Offset to Per Frame content Macro as needed for Enhanced MR and Enhanced CT objects. | W. Corbijn |  |  |
| 1196 | Add phantom-specific Total DLP Values to allow for Head and Body in same RDSR scope of accumulation | D. Clunie |  |  |
| 1203 | PDR Pulse Details in RT Brachy Session Record | U. Busch |  |  |
| 1207 | Change Multi-frame Dimension module usage to “User Option” for IVOCT IOD | B. Revet | Tom Probasco-WG-06 needs more information and examples for when it is meaningful [BR to contact TP] |  |
| 1208 | Clarify the unambigous use of Private Creator Data | B. Revet |  |  |
| 1209 | Incorrect use of private creator tag in PS 3.2 Annex B | B. Revet |  |  |
| 1219 | Clarify Completion Flag for Dose SR | H. Solomon |  |  |
| 1220 | Editorial Corrections for PS3.16 | H. Solomon |  |  |
| 1224 |  | K. O’Donnell |  |  |
| 1225 | Add Procedure Step Completion Codes to UPS | K. O’Donnell |  |  |
| 1226 | Add more specific post-processing codes to CID 9231 | K. O’Donnell |  |  |
| 1227 | Fix Estimated Dose Saving attribute | K. O’Donnell |  |  |

# New Correction Proposals Submitted– assigned CP Number and Editor

The following new change proposals received CP numbers and will be placed in the “Assigned” folder of the next  WG-06 meeting:

| CP# assigned | New Correction Proposals -- Title | Assigned to | Discussion | File ID |
| --- | --- | --- | --- | --- |
| 1228 | Correct inconsistency between Table C.30.3-1 and Table CC.2.5-3 of UPS | K ODonnell |  | Cp_Correct… |
| 1229 | Add glucose to PET acquisition context | D Clunie |  | Dac187 |
| 1230 | Correct terms for fore and hind navicular bone | D Clunie |  | Dac188 |
| 1231 | Correct abbreviation for left in quadruped orientation example | D Clunie |  | Dac189 |
| 1232 | Add iterative reconstruction to CT RDSR | D Clunie |  | Dac190 |
| 1233 | Add Patient Size (Height) to Modality Worklist | B Revet |  | Djh AddSize.. |
| 1234 | Add GBK and GB2312 Character Sets for Chinese Text Encoding | R Horn | In Voting Package August 2012 | Cp_GB2312 |
| 1235 | Inconsistent units of Acquisition Duration (0018,9073) | B Revet | Members present preferred seconds | cpxxEnhUSVol |
| 1236 | Directional flow velocity issues with Enhanced Ultrasound Volume IOD | B Revet |  | cpxxEnhUSVol direct |
| 1237 | Equally spaced Planes in Enhanced US Volume | B Revet |  | cpxxEnhUS VolSpac.. |
| 1238 | Inappropriate for Mandatory Synchronization Module in Enhanced US Volume | B Revet |  | cpxxEnhUS Volsynch. |
| 1239 | Change Value Multiplicity of the Findings container for Pelvic Vascular Structure | B Revet |  | cpxxPelvicVasc |
| 1240 | Replacement of the full organization name for acronym “JIRA” | B Revet |  | Cp_replaceJIRA |
| 1241 | Restrict claming UPS Workitems assigned to another AE | K ODonnell |  | Cp_KOD_restrict |
| 1242 | Clarify that Dimension Index Values start from 1 | D Clunie |  | Dac191 |
| 1243 | Common Instance Reference Module in Image IODs | B Nolte |  | Cp_common inst |

# New Correction Proposals Submitted– No CP Number Assigned  
None

Note:  The comprehensive list of all CPs is posted in ftp://medical.nema.org/MEDICAL/Private/Dicom/WORKGRPS/Wg06/2012/2012-06-11

(enter password.)

# Supplement 154 – Optical Surface Scanner Storage SOP {WI 2010-04-F}

Phillip Liebmann presented the Supplement (sup154_pc_2012-05-28.docx.)  WG-06 made numerous editorial revisions to improve clarity of the text.  The revisions were reflected in the  document posted at the end of the sessionat WG-06: (sup154_pc_2012-06-12.docx.)

The discussion included the review of the comments received in the Public Comment Phase-

The Supplement was approved for Letter Ballot.

# Supplement 147 – Second Generation Radiotherapy {WI 2007-06-B}

C. Schadt first discussed the early software experimentation by WG-07 with the draft supplement (addition to the DICOM standard.)  This was previously referred to by WG-07 as trial implementation.  (Trial implementation makes use of use of the supplement at a later phase of development.)  Because this supplement is very complex and large, WG-07 is undertaking an early experimentation with private data components (private tags and private IODs), prior to the public comment period.  The experience gained will be used by WG-07 (a) to revise the supplement, if needed, (b) to provide guidance for migrating implementations from 1st generation RT, (c) to provide guidance to help IHE RO create an profile that provides implementation guidance to avoid differing interpretations by implementers of this complex supplement.   All on the WG-07 mailing list (members and observers) are invited to perform and report on such experimentation.  The Vice Chair’s company will produce and make available a white paper and possibly open source software to share such experiences.  Input from DICOM library vendors is generally needed for implementation.  It was discussed that such library vendors are expected to participate or monitor the DICOM Standards Committee and/or WG-06.  Accordingly WG-06 advised to create a dictionary of the private tags used in the experimentation in public format and to include the mailing lists of the latter two committees in the distribution of information and invite all to join the WG-07 mailing list for further communication.

Next, C. Schadt presented the Supplement (Version 35, Draft 3.)   This draft included changes made by WG-07 in response to the WG-06 comments made during the March presentation to WG-06.  WG-06 requested numerous revisions. Next Part 17 of Supplement 147 waqsw reviewed started and the explanation of the new IODs and the corresponding information model diagrams were discussed.

After the presentation and after C. Schadt  will finish making the changes asked by WG-06, the document will be  saved to the WG-06 meeting folder as Version 35, Draft 4.

The Supplement will return to WG-06.

# Supplement 160 – Patient Positioning Workflow  {WI 2011-12-A}

C. Schadt provided the background for this supplement.  In short, this has been split off from the original Sup 147 to make the latter smaller.  A new PWP Subgroup has met, in January, for the first time with attendance from 15+ vendors.

This new Supplement of WG-07 on Patient Positioning and Workflow received Supplement Number 160, pursuant to DICOM Work Item 2011-12-A.

The Supplement will return to WG-06.

# Supplement 124 – Communication of Display Parameters {WI 2004-12-B}

Norihisha Matsui and Masatoshi Kanzaki of  WG-11 were again connected from Japan; WG-06  representative K. O’Donnell again assisted in the continued review of Sup. 124.   The detailed review of the Supplement continued for a total of four hours.  Numerous revisions were made.

The Group reviewed terminology and references to AAPM TG-18.

The Supplement will return to WG-06.

# Supplement 158 – Retire General Purpose Worklist {WI 2011-04-B}

Kevin O’Donnell presented for Second Read the draft Supplement.  After a thorough review and line-by-line reading the Group approved it for Public Comment.

# Supplement 161 – Introduction of Web Services in DICOM {WI 2008-04-B}

Tim Culp presented background information and the draft supplement.  This was a First Reading.  This Supplement defines RESTful Services for providing DICOM images and other persistent objects to an Electronic Medical Record/Electronic Health Record (EMR/EHR) system.  Access to DICOM normalized objects is not defined within this supplement.

This supplement deals only with retrieval, corresponding to the evolution of the existing WADO to RESTful Services. Query and notification mechanisms are not defined within this supplement.  Native DICOM can be retrieved, as well as binary data, pixel data, or metadata of the object without the image pixels without any bulk data.

Security is beyond the scope of the RESTful services defined in this supplement.  However the proposed mechanism is fully compatible with generic Web security mechanisms and several security programming recipes are provided for reference.

The Supplement was assigned the number 161.

The Supplement will return to WG-06.

# Supplement 121 – Modality Procedure Plan and Protocol Storage {WI 2006-04-E}

The purpose is to produce text for Public Comment.  K. O’Donnell continued to review the Supplement.  There will be meeting of WG-21 in July to continue development.  The topics mentioned included:

- Intent on designate body part in view

- Distinguish between image that contains e.g. an organ vs. an image of the organ

# Supplement XXX  STOW (Store over the web)

Tim Culp of WG-27 presented a ppt outline of the Supplement: storing and updating DICOM studies over web-based protocol.  WG-06 suggested limiting the resources to RESTful resources so this work item is not held up by the Web Services. The work on STOW-WS API is still to be assigned..

WG-06 will examine additional operations and will advise.

Planned next steps:

1st read: Aug WG-06

PC : Nov. WG-06

T. Culp will define use cases by the August 2012 WG-06 meeting.

# The Supplement will return to WG-06.

- Future Meetings

The Group’s future meeting dates and locations are as follows:

- T-con, prior to the Aug. 2012 meeting – at the call of the Chair

- Meeting August 27 – 31, 2012  in Arlington VA

- T-con, prior to the Nov. 2012 meeting - – at the call of the Chair

- Meeting November 5 – 9, 2012 in Arlington, VA

2013

- Meeting Feb. 4-8, 2013 in Arlington, VA

- T-con, prior to the Feb. 2013  meeting – – at the call of the Chair

- Meeting Apr. 8-12, 2013 in Arlington, VA

- T-con, prior to the Apr. meeting – – at the call of the Chair

- Meeting June 17-21, 2013 in Arlington, VA

- T-con, prior to the June meeting - – at the call of the Chair

- Meeting Aug. 12-16, 2013 in Stockholm, Sweden

- T-con, prior to the Aug meeting – – at the call of the Chair

- Meeting Oct. 14-18, 2013 in Arlington, VA

- T-con, prior to the Oct meeting – at the call of the Chair

- Next Meeting Agenda

A detailed agenda for the next meeting (2012 Aug.) will be prepared at a later date.  The preliminary agenda plan was prepared and recorded in the Agenda Plan for August 2012, which may be found in the June  2012 meeting folder.

- Adjournment

The meeting recessed at 17.30 daily and resumed at 08.30 the next day.  The meeting was adjourned at 17.30 on June 12, 2012.

Reported by:							Stephen Vastagh, Secretary

2012-08-26

Reviewed by counsel: CRS  2012-08-26

Editorial changes, corrected titles of CPs 1236, 1237, 1238