|  | 1300 North 17th Street, Suite 1752 Arlington, VA 22209, USA +1-703- 841-3281 http://dicom.nema.org E-mail: dicom@medicalimaging.org |
| --- | --- |

Meeting MINUTES

###### DICOM WORKING GROUP SIX

(Base Standard)

Aug. 27-31, 2012

### MITA HQ, Arlington, VA

| Members Present | Represented by |  | M | T | W | Th | F |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Agfa Healthcare | Robert Horn |  | x | x | x | x | x |
| CoreLab | David Clunie |  | P | x | x | x | x |
| DatCard Systems | Jere Darling |  |  | x | x | x | x |
| GE Healthcare | Andrei Leontiev |  | x | x | x | x | x |
| JIRA | Makoto Suzuki |  |  | x | x | x |  |
| Konica Minolta Med. Imaging | Michael Laconti |  | x | x | x | x | x |
| Philips Healthcare | Bas Revet |  | x | x | x | x | x |
| Siemens Healthcare | Bjorn Nolte |  | x | x | x | x | x |
| Toshiba Medical Research Institute USA | Kevin O’Donnell |  | x | x | x | x | x |
| Members Absent | Voting Representative |  |  |  |  |  |  |
| * ICCAS (Innovation Center Computer Assisted Surgery (ICCAS)) | Burgert, Oliver |  |  |  |  |  |  |
| *Carl Zeiss Medical Software | Tobias Kurzke |  |  |  |  |  |  |
| *DeJarnette Research Systems | Fred Ganong |  |  |  |  |  |  |
| *Merge Healthcare | Ronald Cowan |  |  |  |  |  |  |
| *Societe Francaise de Radiologie | Joel Chabriais |  |  |  |  |  |  |

| Alternate Voting Representatives, Observers, Others Present | Affiliation |  | M | T | W | T | F |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Balasubramenianm Shridar | Contrast Ad-Hoc/Bayer |  | x | x |  |  |  |
| Baumann, Steffen | Contrast Ad-Hoc /Covidian |  | x | x | x |  |  |
| Busch, Ulrich | WG-07/Varian |  |  | P |  | P |  |
| Costa-Teixeira, Jose | AGFA |  | P | P |  |  |  |
| Culp, Timothy | WG-27/Harris Corp. |  |  |  | x | x | x |
| Fauquex, Jacques | Open Dicom |  |  |  |  | P |  |
| Heaney, David | McKesson |  | P |  | P | P |  |
| Kalafut, John | Contrast Ad-Hoc/Bayer |  | x | x | P |  |  |
| Kantchev, Todd | Salvesta |  | P |  |  |  |  |
| Kanzaki, Masatoshi | WG-11-Real Vision |  |  |  | P |  |  |
| Liebmann, Philipp | WG-24-ICCAS |  | P |  |  |  |  |
| Lingley, William | Merge Healthcare |  |  |  |  | P |  |
| Matsui, Norihisa | WG-11-Shimadzu |  |  |  | P |  |  |
| Matthijssen, Henri | WG-16/Philips |  |  | P | P |  |  |
| Philbin, James | WG-27/Johns Hopkins |  |  |  | x | P | x |
| Pohlhammer, Jeffrey | WG-03/Philips |  | x | x |  |  |  |
| Radko, Radev | Contrast Ad-Hoc/Medtron |  | x | x |  |  |  |
| Smith, Charles | WG-03/NUMA, Inc. |  | x | x |  |  |  |
| Stephen Vastagh | MITA Staff |  | x | x | X | x | x |
| van Dijken, Henk | WG-07/Elekta |  |  | P |  |  |  |
| Whitby, Jonathan | Toshiba |  |  |  | x | x | x |

(* = Does not count toward a quorum,  P = Phone and/or web,  X = Present)

Presiding Officer:							Bas Revet, Chair

# Opening

- The Chair called the meeting to order at 09.30 on August 27, 2012.  The participants identified themselves and their employers.  The secretary reported that a quorum is present.  Members approved the agenda with a change to swap CPs from Tuesday with Sub 157 from Monday.  The minutes of the June 2012 meeting will be approved at the next meeting to allow more time for review.  The Secretary reviewed the antitrust rules.

# Report on Ballots

Committee Secretary Stephen Vastagh reported on the ballots distributed following the June 2012 meeting of WG-06.  The ballots passed.

| ## BALLOT | ## APPROVE | ## DISAPPROVE | ## ABSTAIN | COMMENTS |
| --- | --- | --- | --- | --- |
| CPack 67 | - 25 | 0 | 9 | 6 |
| Sup 154 | - 19 | 0 | 5 | 3 |

# Discussion: CPs in last Letter Ballot CPack-67  - Review  of the Comments, Disposition

- The CPs in the Letter Ballot CPack-67 are listed below together with discussion of comments and decision on Final Text or other disposition.

| CP | Name |  | Consideration of Comments and Other Discussion | Dis-position |
| --- | --- | --- | --- | --- |
| 739 | Add View Code Sequence to CT, MR, SC for Cardiac Views (moved to assigned for coordination with other WGs) | D. Clunie | No comments | FT |
| 1137 | Structure Set Relationship Sequence | U Busch | No comments | FT |
| 1138 | Average Beam Dose Parameter | U Busch | Comments accepted. | FT |
| 1195 |  | B Revet | Comments accepted. | FT |
| 1197 | Unspecified code should be forbidden in extensible context groups for mandatory content items | D Clunie | No comments | FT |
| 1198 | Clarify encoding of Surface Segmentation Point Indices | D Clunie | No comments | FT |
| 1199 | Add Modifier to Processing Parameters Sequence | D Clunie | No comments | FT |
| 1202 | Correct code for Liver | B Revet | No comments | FT |
| 1204 | Identify specific Source used in RT Plan | U Busch | Comments accepted. | FT |
| 1205 | Device Parameters in CT Images | U Busch | Comment on enhanced object was acknowledged, WG-06 intends to address the issue at a later date. Waiting for response from U. Busch in response to Toshiba comments. Decided on name Radiotherapy Equipment Positioning Correlation Macro; thereafter approved for FT. | FT |
| 1206 | Dose Summation Type Clarification | U Busch | No comments | FT |

- WG-06 decided unanimously to approve, as final text, the CPs listed above.

# Discussion of CPs in the Voting Package for Aug 2012

- The CPs listed below have either been proposed for the Aug VP at the past meetings or have been presented for VP beeWG-06 reviewed and discussed the CPs that were in the VP (Voting Package) coming into this meeting.  The CPs were approved for inclusion in Letter Ballot (CPack 68) or directed for further work or review as shown below:

| CP# | Voting Package for Aug. 2012 -- Title | Assigned to | Discussion | Disposition |
| --- | --- | --- | --- | --- |
| 1112 | Add SR Structure for ROI Statistics | H. Solomon | Reviewed numerous changes by D. Clunie. Approved for LB | LB |
| 1152 | Additional document titles for Key Object Selection Document | D Clunie | Approved for LB after substantial review but with no changes. | LB |
| 1200 | Add attributes to Surface Mesh Module | U. Busch | Approved for LB | LB |
| 1201 |  | B Nolte – will involve H Blendinger as Wg-28 member | There was a strong suggestion to extend the description with a example of the workflow and where the value specified play a role. That is in the rationale of the CP but the text in the description is missing this background. Kevin will provide the description. B. Revet noted that WG-02 can also work on it at their October meeting with WG-28 representatives involved in WG-02. Approved for VP provided the listed additional work is completed. of the November meeting. | Nov VP |
| 1210 | Add Modifier to Processing Parameters Sequence | D. Clunie | Approved for LB | LB |
| 1211 | Inconsistent Mean Point Distance in Surface Segmentation | D. Clunie | Approved for LB | LB |
| 1212 | SR IOD with SCOORD3D | D. Clunie | Discussed whether CP is the proper path for a new SOP Class. D. Clunie suggested sending the CP for Public Comment, R. Horn argued that the DICOM Community is not conditioned to pay as much attention to CPs as to Supplements. Harry suggested before that this is covered under the Evidence Document Work Item. ACTION: D. Clunie will draft a Supplement. Sup. 162 assigned: “Comprehensive 3D SR Storage SOP Class” under [WI 2003-12B] | CANCEL CP Reissue as Supplement 162 |
| 1213 | Clarify homogeneous transformation matrix constraints | D. Clunie | Approved for LB | LB |
| 1214 | Add Tracking Identifier Template to Measurement Templates | D. Clunie | Approved for LB | LB |
| 1215 | Require GSPS for Enhanced PET Storage for consistency with other enhanced images | D. Clunie | Approved for LB | LB |
| 1216 | Instance Creation Date and Time not changed on coercion | D. Clunie | Approved for LB | LB |
| 1218 | Correct further errors in UCUM cades and related meanings | B Nolte | Approved for LB | LB |
| 1221 | Add note on VR conflict in data dictionary | D. Clunie | Approved for LB | LB |
| 1222 | Amend DIMSE Service status codes | B. Revet | Approved for LB | LB |
| 1223 |  | B. Revet | Discussed and decided that it would be desirable to organize a t-con between DICOM WG-06 and the IEC Team (H Blendinger, Chair) (and of course WG-02 and WG-28) to discuss the terminology used. It is considered important that this is harmonized between IEC and DICOM. Doing this during CP sessions will not result is a quick result (with B. Revet as messenger) and will slow down the process. A separate full time session is much more productive. Secretary’s Note: The Chair of the IEC TF agrees and offers that the IEC TF meets in Paris immediately preceding the WG-02 mtg. The goal of the IEC TF is to stabilize the current IEC 61910-1 draft to a CDV (Committee Draft for Voting). This will start the IEC process to finally have a Standard. So to be able to create the CDV, we need the IEC 61910-1 Draft and CP 1223 to be harmonized. Accordingly the t-con needs to occur before the October IEC meeting. | Nov VP |
| 1234 | Add GBK and GB2312 Character Sets for Chinese Text Encoding | R Horn | The CP would allow duplication of characters that is undesirable in general. However, it was pointed out that such duplicate use is in effect now and disallowing it in DICOM would not change that fact. Approved for LB after removing all changes except adding GB2312 and GKB to the standard. | LB |

Secretary’s note:

Upon canceling CP-1212 and the decision to submit as supplement the Supplement Number 162 has been assigned.  Subsequently, D. Clunie drafted and posted in this meeting’s folder draft Supplement 162: Comprehensive 3D SR Storage SOP Class.

# Letter Ballot CPack-68

- There were no additional CPs approved for letter ballot other than those from the August Voting Packet.

WG-06 decided unanimously to submit for legal review and thereafter circulate in a Letter Ballot CPack 68, for vote by the DICOM Standards Committee, the Correction Proposals (CPs) listed below:

| CP# | Letter Ballot to Be Issued CPack-68 – Title |  |  |  |
| --- | --- | --- | --- | --- |
| 1112 | Add SR Structure for ROI Statistics |  |  |  |
| 1152 | Additional document titles for Key Object Selection Document |  |  |  |
| 1200 | Add attributes to Surface Mesh Module |  |  |  |
| 1210 | Add Modifier to Processing Parameters Sequence |  |  |  |
| 1211 | Inconsistent Mean Point Distance in Surface Segmentation |  |  |  |
| 1213 | Clarify homogeneous transformation matrix constraints |  |  |  |
| 1214 | Add Tracking Identifier Template to Measurement Templates |  |  |  |
| 1215 | Require GSPS for Enhanced PET Storage for consistency with other enhanced images |  |  |  |
| 1216 | Instance Creation Date and Time not changed on coercion |  |  |  |
| 1218 | Correct further errors in UCUM cades and related meanings |  |  |  |
| 1221 | Add note on VR conflict in data dictionary |  |  |  |
| 1222 | Amend DIMSE Service status codes |  |  |  |
| 1234 | Add GBK and GB2312 Character Sets for Chinese Text Encoding |  |  |  |

# Discussion of assigned CPs (CPs in the “Assigned” meeting folder and other previously assigned CPs)

(The table below records discussion of CPs that were NOT  in the last Letter Ballot CPack 67 (Table 3) NOR in the current (Aug 2012) Voting Packet (Table 4.)

| CP# | Title | Assigned to | Discussion | Dis-position |
| --- | --- | --- | --- | --- |
| 1196 | Add phantom-specific Total DLP Values to allow for Head and Body in same RDSR scope of accumulation | D. Clunie | Need to be discussed in WG-21 | Remains assigned |
| 1208 | Clarify the unambigous use of Private Creator Data | B. Revet | B Revet clarified that the intent of the CP is to prevent the same identifier issued for different blocks. It was suggested to state that the “same privet creator shall not be used for different blocks.” ACTION: B Revet will reword it to include more normative text and include examples in the notes. Also make it clear that there is no implication in the order of the blocks. | Remains assigned |
| 1209 | Incorrect use of private creator tag in PS 3.2 Annex B | B. Revet | Approved for VP | Nov VP |
| 1217 | Add Extensible SR SOP Class for new types of Content Item | D.. Clunie | June discussion: Postponed discussion when D Clunie and R Horn are present in Aug ++++++++++++++++++++++++++++++ August: should there also be a Supplement developed instead of a CP? Discussed what warning might need to be given if content is not understood. The question was posed whether this should be pursued. It was noted that SR cannot handle many content such as vectors (e.g. BP waveform.) Members were concerned of safety but acknowledged the value of this CP. Another value is the advancement of SR in general. This CP could add contents of the Evidence Document. It was suggested to include a few illustrations. | Remains assigned. |
| 1229 | Add glucose to PET acquisition context | D Clunie | Approved for Voting Package November 2012 with no changes | Nov VP |
| 1230 | Correct terms for fore and hind navicular bone | D Clunie | Approved for Voting Package November 2012 with no changes | Nov VP |
| 1231 | Correct abbreviation for left in quadruped orientation example | D Clunie | Approved for Voting Package November 2012 with no changes | Nov VP |
| 1232 | Add iterative reconstruction to CT RDSR | D Clunie | Approved for Voting Package November 2012 with no changes | Nov VP |
| 1233 | Add Patient Size (Height) to Modality Worklist | B Revet | Approved for Voting Package November 2012 with no changes | Nov VP |
| 1235 | Inconsistent units of Acquisition Duration (0018,9073) | B Revet | Members present preferred seconds. Approved for Voting Package November 2012 with no changes | Nov VP |
| 1236 | Directional flow velocity issues with Enhanced Ultrasound Volume IOD | B Revet | Approved for Voting Package November 2012 after changing the new attribute name into Pixel Data Value at Which Velocity is Zero | Nov VP |
| 1237 | Equally spaced Planes in Enhanced US Volume | B Revet | Approved for Voting Package November 2012 after revising the addition to the attribute description of Position Measuring Device Used | Nov VP |
| 1238 | Inappropriate for Mandatory Synchronization Module in Enhanced US Volume | B Revet | Cancelled after considering that impact of having always Synchronization Module is very small (a single UID) | Cancelled |
| 1239 | Change Value Multiplicity of the Findings container for Pelvic Vascular Structure | B Revet | Approved for Voting Package November 2012 with no changes | Nov VP |
| 1240 | Replacement of the full organization name for acronym “JIRA” | B Revet | Approved for Voting Package November 2012 with no changes | Nov VP |
| 1242 | Clarify that Dimension Index Values start from 1 | D Clunie | Approved for Voting Package November 2012 with no changes | Nov VP |
| 1243 | Common Instance Reference Module in Image IODs | B Nolte | Approved for Voting Package November 2012 with no changes | Nov VP |

# VP (Voting Package) for Nov 2012

- Based on discussions and decisions as recorded in Tables 3, 4, and/or 6, the Voting package for August 2012 is:  (CPs are copied from previous tables for convenient reference.)

| CP# | Title | Assigned to |  |  |
| --- | --- | --- | --- | --- |
| 1201 |  | B Nolte – will involve H Blendinger as Wg-28 member |  |  |
| 1209 | Incorrect use of private creator tag in PS 3.2 Annex B | B. Revet |  |  |
| 1223 |  | B. Revet |  |  |
| 1229 | Add glucose to PET acquisition context | D Clunie |  |  |
| 1230 | Correct terms for fore and hind navicular bone | D Clunie |  |  |
| 1231 | Correct abbreviation for left in quadruped orientation example | D Clunie |  |  |
| 1232 | Add iterative reconstruction to CT RDSR | D Clunie |  |  |
| 1233 | Add Patient Size (Height) to Modality Worklist | B Revet |  |  |
| 1235 | Inconsistent units of Acquisition Duration (0018,9073) | B Revet |  |  |
| 1236 | Inappropriate for Mandatory Synchronization Module in Enhanced US Volume | B Revet |  |  |
| 1237 | Directional flow velocity issues with Enhanced Ultrasound Volume IOD | B Revet |  |  |
| 1239 | Change Value Multiplicity of the Findings container for Pelvic Vascular Structure | B Revet |  |  |
| 1240 | Replacement of the full organization name for acronym “JIRA” | B Revet |  |  |
| 1242 | Clarify that Dimension Index Values start from 1 | D Clunie |  |  |
| 1243 | Common Instance Reference Module in Image IODs | B Nolte |  |  |

# List of Assigned CPs that were discussed at this meeting and remain “Assigned”
As recorded in Tables 4, 5, and/or 7 (CPs are copied from previous tables for convenient reference):

| CP# | Title | Assigned to | Discussion | Dis-position |
| --- | --- | --- | --- | --- |
| 1196 | Add phantom-specific Total DLP Values to allow for Head and Body in same RDSR scope of accumulation | D. Clunie | Need to be discussed in WG-21 | Remains assigned |
| 1208 | Clarify the unambigous use of Private Creator Data | B. Revet | B Revet clarified that the intent of the CP is to prevent the same identifier issued for different blocks. It was suggested to state that the “same privet creator shall not be used for different blocks.” ACTION: B Revet will reword it to include more normative text and include examples in the notes. Also make it clear that there is no implication in the order of the blocks. | Remains assigned |
| 1217 | Add Extensible SR SOP Class for new types of Content Item | D.. Clunie | June discussion: Postponed discussion when D Clunie and R Horn are present in Aug ++++++++++++++++++++++++++++++ Aug discussion: should there also be a Supplement developed instead of a CP? Discussed what warning might need to be given if content is not understood. The question was posed whether this should be pursued. It was noted that SR cannot handle many content such as vectors (e.g. BP waveform.) Members were concern of safety but acknowledged the value of this CP. Another value is the advancement of SR in general. This CP could add contents of the Evidence Document. It was suggested to include a few illustrations. | Remains assigned. |

- List of assigned CPs that were NOT discussed at this meeting and remain “Assigned”
The following, previously assigned, correction proposals were not discussed at this meeting and remain “Assigned”.

| CP# | Title | Assigned to | Discussion |  |
| --- | --- | --- | --- | --- |
|  | LIST AT END OF JUNE MTG |  |  |  |
| 145 | Curve Data Value Representation Clarification | D. Clunie |  |  |
| 235 | Clarify NM detector start angle usage | J. Pohlhammer |  |  |
| 296 | Eliminate Duplicate Terms and Provide Missing Definitions | H. Solomon |  |  |
| 375 | - Correct CID 4017 – Intra-oral Anatomy Modifier | K. O’Donnell |  |  |
| 430 | Additions to PS 3.16 for Cardiac Cath SR | H. Solomon |  |  |
| 450 | Correct Myocardium Mass | D. Sluis |  |  |
| 464 | Add SOP Common to the Print-related N-CREATEs | R. Horn |  |  |
| 471 | Add Common Finding Template | H. Solomon |  |  |
| 591 | Attribute Identifier List for Errors in Sequence Attributes | H. Solomon |  |  |
| 596 | Add Request and Schedule Information to Results and MPPS | R. Horn |  |  |
| 597 | Instance Availability Notification Structure | H. Solomon |  |  |
| 757 | Add RTSS ROI Feature UID | H. Solomon |  |  |
| 802 | Incorporate Fixes on 702 and 706 | D. Clunie |  |  |
| 804 | Indicator for Cone Beam CT | K. O’Donnell |  |  |
| 811 | Note for Forwarders Regarding SCP-assigned Instance UIDs during N-CREATE | D. Harvey |  |  |
| 812 | Clarification that Pixel Padding Value Is a Raw Value before Any Transformations | D. Harvey |  |  |
| 833 | Transfer Syntax Availability and Selection for Q/R | R. Horn |  |  |
| 836 | Clarify Usage of Attributes for Normalized IODs | K. O’Donnell |  |  |
| 838 | Consistent Naming of Error Code Tables | K. O’Donnell |  |  |
| 934 | Fix Inconsistency of C-FIND Only Number of Patient Related XXX Attributes | D. Harvey |  |  |
| 941 | More Clarification of Instance References | D. Clunie |  |  |
| 991 | Clarify Handling of Private Data in the Retrieve Without Bulk Data Service | R. Horn |  |  |
| 992 | Clarify Allowable Length Values for Certain PDUs | K. O’ Donnell |  |  |
| 1005 | Make Keywords Valid Java/C Identifiers | R. Horn |  |  |
| 1006 | Add Keywords to Command Dictionary | R. Horn |  |  |
| 1013 | RT Ion Beams Recording Clarifications | WG-07 |  |  |
| 1019 | Add Body Part at the Study Level | K O’Donnell |  |  |
| 1030 | Re-factor Common Attributes for Enhanced Images | B. Revet |  |  |
| 1031 | Use of OIDs and Non-ASCI Characters in Codes | R. Horn |  |  |
| 1032 | Fix Position Angle Sign Inconsistency in Mammography Image | K. O’Donnell |  |  |
| 1066 | Encoding of Attributes with Value Length > 64KiB with Explicit VR | R. Horn |  |  |
| 1086 | New DICOMDIR ZIP File Extension | R. Horn |  |  |
| 1106 | Assign a Namespace for Audit Schema Elements | R. Horn |  |  |
| 1115 | Add Dose Reference Point Codes for CR/DR | K. O’Donnell |  |  |
| 1126 | Corrections LOINC Codes Sup78 | B. Revet |  |  |
| 1127 | Add Fields for Organ Dose to Dose SR | K. O’Donnell |  |  |
| 1163 | Conformance for Application Hosting | H Solomon |  |  |
| 1164 | Coronary Artery Codes | H Solomon |  |  |
| 1173 | Clarify basis for Total Fluoro Time and Total Acquisition Time in XA Dose Report | K O’Donnell |  |  |
| 1180 | Use LOINC_Short Name | H. Solomon | Was in Aug 12 VP |  |
| 1189 | Add usage of Temporal Position Time Offset to Per Frame content Macro as needed for Enhanced MR and Enhanced CT objects. | W. Corbijn |  |  |
| 1203 | PDR Pulse Details in RT Brachy Session Record | U. Busch |  |  |
| 1207 | Change Multi-frame Dimension module usage to “User Option” for IVOCT IOD | B. Revet | Tom Probasco-WG-06 needs more information and examples for when it is meaningful [BR to contact TP] |  |
| 1219 | Clarify Completion Flag for Dose SR | H. Solomon |  |  |
| 1220 | Editorial Corrections for PS3.16 | H. Solomon |  |  |
| 1228 | Correct inconsistency between Table C.30.3-1 and Table CC.2.5-3 of UPS | K O’Donnell |  |  |
| 1241 | Restrict claming UPS Workitems assigned to another AE | K O’Donnell |  |  |

# New CPs Submitted that were assigned a CP Number and an Editor

The following new change proposals were received and posted in the “New” subdirectory of the meeting folder.  CP numbers shown were assigned by WG-06.  They will be placed in the “Assigned” folder of the next  WG-06 meeting:

| CP# assigned | New Correction Proposals -- Title | Assigned to | Discussion | File ID |
| --- | --- | --- | --- | --- |
| 1244 |  | U Busch |  | RT22_Additional… |
| 1245 |  | U Busch | U Busch proposed to retire this SOP Class. Many references were found in the Standard. WG-06 suggested adding the note to all occurrences. U Busch will re-write the CP and add note and submit to the Nov. Mtg. | RT_23 NoteAbout…… |
| 1246 |  | U Busch |  | RT_24_BeamAuto |
| 1247 |  | U Busch |  | RT_25_Rename CID 21 |
| 1248 |  | U Busch | Aug. 2012: The CP will be completed for the next meeting. | RT_26_Indication… |
| 1249 |  | U Busch | Aug. 2012: This CP is the RT-relevant response to the concept of CP-1213. D. Clunie noted that RT Structure Set has been implemented outside of RT community and pushback is expected. WG-06 expressed doubts about removing something that works. RT hopes that segmentation IOD will replace it. Expect argument to allow use the registration object but must note that usage outside RT is permitted. Deprecating individual attributes concerns WG-06. | RT_27-Deprecation … |
| 1250 | Stereotactic Information for Breast X-Ray | B. Revet |  | Cpp__breast XRay |
| 1251 |  | D. Clunie |  | Dac 193 |
| 1252 |  | D. Clunie |  | Dac 194 |
| 1253 |  | B. Revet |  | hb-carify… |
| 1254 |  | B. Revet |  | hb_Correct.. |
| 1255 |  | U Busch | U Busch noted that this is an editorial correction only. | RT_28_Incorrect Note,,,, |

# New Correction Proposals Submitted– No CP Number Assigned

None

Note:  The comprehensive updated list of all CPs is posted in this meetings folder.

# Supplement 154 – Optical Surface Scanner Storage SOP {WI 2010-04-F}

Phillip Liebmann presented the Supplement with changes made as a result of the public comment s received.  The public comments were accepted and adopted.

Next steps: combine figures, add reference, directory codes, revise figure.

WG-06 will review these last revisions when completed.

The Supplement will be reviewed again for Final Text.

Supplement 159 – NM Dose Structured Report {WI 2010-04-B}

This was a reading for achieving Public Comment.

C. Smith presented the draft Supplement <sup159_02.doc> which was previously posted in the meeting folder.

Following the first session, the changes recommended by WG-06 were made and the document Sup159_03.doc and Sup_159_WIP03.doc were posted in the mtg folder.

Additional discussion took place in Session two on 8/28 and further changes will be made by Editor C. Smith and an updated draft < Sup159_04 3.doc.

In-depth technical discussions were conducted on several sections of the draft Supplement during both sessions.  The chief topics discussed included:

- Open issue 1 – clarify it is Time Synch – go ahead and add time synch and note here that we have done that. Leave the item open for comments,

- Open issue 2 – describe that the SR is the definitive source of the dose information – it was historically required in the images – now the SR is replacing this.

- Add a closed issue that this supplement does not describe how to combine or comparing x-ray and administered dose into a whole rolled up report.

- Need to revise all measurements to add a method as a sub-modifier, also including time and date of the measurement.

- Many small editorial changes were made.

- Leave the “Give- Sub-ID” issue open. Describe what we are trying to achieve and why. We are trying to get the prescription number. Do we need this other stuff?

- Should use anatomical codes for approach site.

- Open question: should there be a separate IOD and SOP Class for this, or should it re-use the existingr CT Dose SR (with an additional template)? The devices are different.

- Open issue: Which codes and coding scheme designators for NDC and RXNORM.

- Open issue: how much detail needs to go into estimates of fetal dose?

- Open issue: Should administered activity be in Bq or MBq? The issue is that NM and PET IODs use different units now.

- Open issue: Do we need to track “traceable to DMS”?

The Supplement will return to WG-06 at the Nov. 5-9, 2012 meeting (Arlington, VA.)

# Supplement XXX – Contrast Injection SR

This was a Second Reading.

J Kalafut presented the concepts and flow of the Supplement using a presentation <SupXXX_ContrastAgentAdministrationReport_WG-06_Aug27_2012_Meet_Updated.docx>.  Thereafter S. Balasubramanian continued the presentation of the detailed components (TID…) of the supplement. In-depth technical discussion was conducted on the concepts for the Supplement and on the components of the future Supplement that were included in the presentation.  The chief topics discussed included:

- The Ad-hoc group chair reviewed the motivations, clinical needs and rationale for a standards-based exchange of contrast administration records. The review included descriptions of contrast usage in diagnostic and interventional radiology and cardiology settings.

- A detailed review of use-cases and workflows generated by the ad hoc group was presented. WG6 members actively participated by asking questions, adding clarifications and requesting further clarifications. WG6 agreed that making distinctions among contrast prescribing, preparation, administration and consideration of varying modalities made sense.

- A review of the data model and data structures as proposed by the ad hoc group was also conducted. The structure was presented so as to assist the WG in deciding between encoding options for the administration record – structure report or attribute.

- A review of other WG inputs (21,3 and 16) were reviewed. It was noted that WG16 in particular preferred an attributed encoding solution to ease implementation burdens on modality vendors.

- An active area of discussion centered on methods to encode the vector-based data in the supplement, most notably pressures and flow rates of contrast media delivered to the patient. Three options were discussed – encoding the data as a DICOM waveform object  and referencing it within the contrast administration record, persisting the data as discrete “stair step” data within an SR or attribute context, and defining a new mechanism to persist vector data within an SR.

- WG6 ultimately directed the AdHoc group to refine the in-process supplement work to encode the overall contrast record as an SR DICOM object but to also have a DICOM attribute object for consumption by modalities during an imaging examination. The DICOM attribute object would primarily consist of contrast “step” data emitted by the Infusion Manager actor (not an official IHE actor) during a contrast-enhanced imaging procedure.

- The AdHoc group will accommodate the mandates of WG6 and will do the following prior to the next WG5 meeting:

  - Update the current pre-Supplement version of the document so that it is based upon DICOM/NEMA templates

  - Update and refine the SR and DICOM objects in a timely fashion for review at the November WG6 meeting. The goal is to then allow for PC sometime post November/early 2013.

WG-06 further suggested that the Contrast Ad-Hoc seeks assistance with the writing of the Supplement from DICOM Consultants to save time.

The Supplement will return to WG-06 for continued reading.

# Supplement 147 – Second Generation Radiotherapy {WI 2007-06-B}

U. Busch presented Sup147-Version 36 Draft 5 – that has been posted in the meeting “Sup” directory of the meeing folder, during the presentation.  This draft included updates and changes made in Sup 147 in response of to the WG-06 review in June 20124 and updates that have been added as a result of further work of WG-07.  The review went through the changes in all sections of the supplement, beside Part 17. The latter one will be revised by WG-07 prior to public comment. It was agreed, that a full reading of that Part 17 is not required for public comment. However, WG-07 may seek advice from WG-06 on specific sections of Part 17 prior to public comment.

Following this session and after U. Busch will finish making the changes as per the discussion in this meeting the document will be saved to the WG-06 meeting folder under as Version 36 (final revision). Subsequent changes will continue in Version 37.

The Supplement will return to WG-06 for continued reading, at the Nov. 5-9, 2012 WG-06 meeting in Arlington, VA.

# Supplement 124 – Communication of Display ParaKEVmeters {WI 2004-12-B}

This was a continued reading to obtain approval for Public Comment.

K. O’Donnell presented the draft Supplement on behalf of  Norihisha Matsui and Masatoshi Kanzaki of  WG-11 who were connected by phone from Japan;   <sup124_18F.doc> which was previously posted in the meeting folder.  Version 18F included the changes made pursuant to the last discussion by WG-06.

Detailed technical discussion was conducted.   Notes on the discussion as well as revisions made have been captured in a new version of the document, which was posted <sup1214_18F+WG06.doc.

The Supplement will return to WG-06 for continued reading.

# Supplement 158 – Retire General Purpose Worklist {WI 2011-04-B}

The goal was to approve for Letter Ballot.

Editor, K. O’Donnell presented the public comments: a Public Comment noted that the broad description “Procedure Step SOP Class”, which was chosen by WG-06, is not defined.  In response, it was decided to add a definition to Part 4 of the Standard.

After a thorough review the Sup 158 was approved it for Letter Ballot.

# New Work Item Proposal-Delete Entire Study

A New Work Item Proposal was discussed by the DSC, “Ability to Delete an Entire Study for Data Retention Purposes”  [2012-06-A]

The Group discussed that the Standard should not allow the deletion of a study without reviewing the content.  Legal considerations were also mentioned. The Group asked the author, Dave Heaney of McKesson, to join in a t-con to discuss the proposal.

Dave Heaney explained that the use case is deletion of studies for data retention purposes, requested .  Such action is implemented where data retention policies require the erasure of all traces of records, although there is generally an audit trail of the erasure.  During the discussion WG-06 suggested to delete the content and if no content, show no results.

The conclusion of the discussion was the unanimous confirmation by WG-06 of the DSC action rejecting the New Work Item Proposal. It is not a valid DICOM operation to delete all information of a study (that may include non-DICOM data). Alternate ways to handle a study delete is deleting all the instances of a study with a KOS message and react to a DICOM query for an empty study as if the study not longer exists. IHE may provide other operations that can handle the deletion of a complete study.

# Supplement 161 – Introduction of Web Services in DICOM {NWIP 2008-04-B}

This was a reading for achieving Public Comment.

T. Culp presented the draft Supplement <sup161_11.1_WADO-RS.doc> which was previously posted in the meeting folder and included the changes directed by WG-06 , during the previous presentation in June 2012 .

Following the first session, further changes were recommended by WG-06 and a revision 11.2 was posted.  Some time became available on the Agenda and there was an opportunity on 8/30 to conduct a second session after which the Sup 161 was approved for Public Comment subject to further changes to be made based on the timdiscussions to be reviewed by the Chair Bas Revet.

The PC version is to be named  <sup161_pc.doc> and sup_161_pc.pdf>

In-depth technical discussions were conducted on several sections of the draft Supplement during both sessions.  The chief topics discussed included:

- The filters suggested at the DICOM WG-06 Meeting in Berlin (pixeldataonly, nonprivatedata, and sopclassuid) were determined to be unnecessary. The resources are flexible enough as defined to allow for returning specific frames or specific bulk data items without needing to add the complexity of supporting multiple parameters.

- Retrieving pixel data frames by “Calculated Frame Lists” was removed since most users are only interested in retrieving frames using a simple frame list. The restriction for the frame numbers to be monotonically increasing was removed. The client can request frames in any order and the server must return frames in the order requested.

- The Transfer Syntax UID is no longer specified in the Accept-Encoding and Content-Encoding header parameters since these non-standard encodings might cause issues with some clients. Instead they are specified as custom headers X-DICOM-Accept TS and X-DICOM-TS, respectively.

- multipart/mixed was changed to multipart/related since it allows you to specify the content type o the first part of the message. E.g. Accept: multipart/related; type=application/dicom.

- application/xml changed to application/dicom+xml.

- application/wado-rs was removed for simplicity. XML+Bulkdata can still be retrieved by making two calls, RetrieveMetata (application/dicom+xml) and RetrieveStudy (application/octet-stream).

- The content headers are now specified in more detail.

- The RetrieveBulkdata resource was changed from specifying study, series, SOP Instance UIDs, and bulk data UUID to just specifying bulk data UUID. Bulk data must be globally unique.

The Supplement was approved for Public Comment.

# Supplement 157 – Multi-Frame Converted Legacy Images {WI 2011-09-A}

This was a reading to obtain approval for Public Comment.

D. Clunie presented the draft Supplement <sup157_08.doc>  which was previously posted in the meeting folder.

Detailed technical discussions were conducted on several sections of the draft Supplement during both sessions.

Following this discussion a new version of the document will be posted to record the revisions made as well as the comments recorded, during the discussion.

The Supplement will return to WG-06 for continued reading.

# Supplement 121 – Modality Procedure Plan and Protocol Storage {WI 2006-04-E}

This was a reading to obtain approval for Public Comment.

The file sup121_11d.doc, which included material from the recent meeting of WG-21 in July 2012 and subsequent changes, was previously posted in the WG-6 meeting folder.

K. O’Donnell presented the posted draft and several subsequent revisions during the course of the meeting.

The discussion included:

- the updated SOP Class names, use case priorities and concepts refined in the recent WG-21 meeting

- review and revision of the Reference Location Macro and associated anatomic codesets

- review and revision of the Assertion Macro and the concept of recording protocol approvals

- the need for patient positioning codes and instructions

Following this discussion a new version of the document will be posted to record the revisions made as well as the comments recorded, during the discussion.

The Supplement will be worked on further in WG-21 with a focus on the contents of the Acquisition and Reconstruction Technique Modules

The Supplement will return to WG-06 for continued reading.

# Supplement 162, Comprehensive 3D SR Storage SOP Class [WI 2003-12B]

The Supplement was developed following the decision by WG-06 that a CP (CP-1212) was not appropriate for creating a new SOP Class.  CP-1212  was canceled and Sup. 161 was approved for Public Comment.

# Supplement 163,  STOW (Store Over the Web) [WI-2011-09-B]

This was a First Reading.

T. Culp presented the draft Supplement <supxxx_02_STOW-RS.doc> which was previously posted in the meeting folder.  This document  included the changes directed by WG-06 , during the previous presentation in June 2012 .

A thorough technical discussion was conducted.  This included a discussion of the use cases and the pros and cons of using the various services.  The chief topics discussed included:

- Much of the discussion centered around whether or not all four resources presented were required.

- The “CheckStudyExists” resource was  questioned since there is not a direct equivalent in the DICOM C-Store use case. If a user wanted to check if a study existed, it would most likely be done through a QIDO Query call as opposed to a STOW RESTful resource. Also, knowing the study exists is not helpful because it doesn’t tell you what SOP Instances are actually stored. It was suggested an efficient HEAD method call on the SOP Instance UID might be more useful than a HEAD method call on the Study Instance UID.

- The CommitInstance resource could take a long time based on the underlying storage. It was not designed to be a synchronous operation. In explanatory information, it could be explained to defer to when a client is ready to delete from their system as opposed to requesting a commitment at transfer time.

- The ServerInfo resource may not be required since it could be assumed all Transfer Syntax UIDs are handled by modern servers. A failure response can always be returned if TS is not recognized).

- The StoreStudy resource was discussed and it was decided to optionally specify the Study Instance UID. If the event it is unspecified, the Study Instance UID is discovered in the data as it sent allowing the user to send multiple studies at the same time (as in the case of importers). If the Study Instance UID is specified, then it implies all the instances are part of the specified study and it will generate an error condition otherwise.

- Bulk data ids and boundaries need not be preserved when storing a study. Round trip may not produce the same results (although no medical content will ever be changed). The client is required to call RetrieveMetadata prior to retrieving any bulk data. The server can choose to restructure the bulk data so that it is more efficient to store/transport.

- Action is needed to map all C-Store responses with HTTP response codes.

The document with the revisions and comments recorded during this meeting has been saved to this  meeting’s folder on the ftp server: <<supxxx_02_STOW-RS.doc>

The Supplement will return to WG-06 at the Nov. 5-9, 2012 meeting..

# Future Meetings

- T-con, prior to the Nov. 2012 meeting - – at the call of the Chair

- Meeting November 5 – 9, 2012 in Arlington, VA

2013

- Meeting Feb. 4-8, 2013 in Arlington, VA

- T-con, prior to the Feb. 2013  meeting – – at the call of the Chair

- Meeting Apr. 8-12, 2013 in Arlington, VA

- T-con, prior to the Apr. meeting – – at the call of the Chair

- Meeting June 17-21, 2013 in Arlington, VA

- T-con, prior to the June meeting - – at the call of the Chair

- Meeting Aug. 12-16, 2013 in Stockholm, Sweden

- T-con, prior to the Aug meeting – – at the call of the Chair

- Meeting Oct. 14-18, 2013 in Arlington, VA

- T-con, prior to the Oct meeting – at the call of the Chair

# Next Meeting Agenda

A detailed agenda for the next meeting (2012 Nov.) will be prepared at a later date.  The preliminary agenda plan was prepared and recorded in the ‘Agenda Plan for Nov. 2012’, which may be found in the Aug.  2012 meeting folder.

# Adjournment

The meeting recessed at 17.30 daily and resumed at 08.30 the next day.  The meeting was adjourned at 14.00 on Aug. 30, 2012.

Reported by:							Stephen Vastagh, Secretary

2012-09-07

Reviewed by counsel: 		CRS 2012-09-08